Pediatric Gait Analysis Video Capture System — Project Summary

1. Overall Aim

To build a low-cost, synchronized, multi-camera video recording system for clinical gait analysis of children, using Raspberry Pi 4 boards and Raspberry Pi Global Shutter Cameras, capable of capturing 60 fps footage with controllable optics and synchronized triggering.

2. System Overview

| Feature | Specification |
| --- | --- |
| Number of cameras | 3 cameras |
| Frame rate | 60 fps per camera |
| Synchronization | Hardware trigger line shared between all cameras (one Pi acts as master trigger) |
| Recording control | libcamera or Picamera2 stack on each Raspberry Pi |
| Use case | Pediatric gait analysis – sagittal, coronal, and detailed distal (foot) views |
| Video storage | Local microSD card or USB 3.0 SSD on each Pi |

3. Camera Configuration

| Camera | Plane / View | Typical distance | Focal length | Purpose |
| --- | --- | --- | --- | --- |
| Cam 1 | Sagittal (side) | 2.5–3 m | 6 mm CS-mount | Full-body sagittal view |
| Cam 2 | Coronal (front/back) | 2.5–3 m | 6–8 mm CS-mount | Full-body coronal view |
| Cam 3 | Coronal (close-up feet) | 2–3 m | 8–50 mm zoom (C-mount) or 16 mm CS | Detailed lower-limb / foot tracking |

Lens characteristics:
All lenses are rectilinear (non-fisheye).

Approximate coverage at 2.5–3 m:

- 6–8 mm → ~2 m vertical field (ideal for children)

- 12–16 mm → ~0.8–1 m vertical field (ideal for foot region)

4. Complete Hardware List

| Component | Qty | Description / Notes | Supplier / Example |
| --- | --- | --- | --- |
| Raspberry Pi 4 Model B (8 GB RAM) | 3 | Main processing boards | Piitel |
| Raspberry Pi Global Shutter Camera (IMX296) | 3 | High-speed global shutter image sensors | Piitel |
| 6 mm CS-mount wide-angle lenses | 2 | Full-body sagittal + coronal | Piitel |
| 8–50 mm zoom (C-mount) or 16 mm CS-mount lens | 1 | Close-up foot view | Piitel |
| Arducam Flex Cable Set (7 pcs, 7–100 cm) | 1 set | Long CSI cables for tripod mounting | Piitel |
| Official Raspberry Pi 4 USB-C Power Supply (5.1 V / 3 A) | 3 | Stable power for each Pi | Piitel |
| MicroSD cards (32–64 GB, Class 10 or UHS-I) | 3 | OS + video storage | Piitel (SanDisk Ultra or High Endurance) |
| Cooling kits (heatsink + fan) | 3 | Prevents thermal throttling | Local (E.g. Telsys / Farnell / Digi-Key Israel) |
| Tripods or light stands | 3 | Stable camera positioning | Photo stores |
| Tripod mount plates for Pi GS camera | 3 | For attaching camera PCBs + lenses to tripods | Online (Arducam or generic HQ brackets) |
| Dupont jumper wires + breadboard | 1 set | Trigger GPIO wiring (master / slave sync) | Local electronics store |
| Continuous LED COB lights (60–100 W, CRI ≥ 95) | 2–3 | Flicker-free illumination for short exposure (1/500–1/1000 s) | Photo suppliers (Nanlite Forza 60B / Amaran 100d S / Godox SL60W) |
| Light stands for LED panels | 2–3 | Adjustable height & angle | Photo suppliers |
| Optional USB 3.0 SSDs | 3 | For long continuous recordings without frame drops | Electronics stores |
| Network switch / router (optional) | 1 | For multi-Pi remote control and data transfer | Generic network hardware |

5. Trigger Synchronization Setup

- Each Global Shutter Camera supports an external trigger input (XTR pin).

- One Pi acts as master trigger, sending a 3.3 V TTL pulse via GPIO to the XTR inputs of all cameras (shared ground).

- Slave Pis receive trigger signals on their XTR pins → simultaneous exposure.

- Configure each camera for external trigger mode in libcamera:

libcamera-vid --framerate 60 --width 1456 --height 1088 --ext-trig ...

6. Software Stack

| Layer | Tool | Function |
| --- | --- | --- |
| Operating system | Raspberry Pi OS (Bookworm 2023 +) | Base Linux environment |
| Camera framework | libcamera | Camera streaming control, exposure, triggers |
| Python API (optional) | Picamera2 | Scripting, automation, and GPIO trigger management |
| Analysis software (PC) | Kinovea | Frame synchronization, measurement, annotation |

7. Lighting Specifications (can be left out if too expensive)

- Continuous LED COB panels rated 60–100 W, CRI ≥ 95, and flicker-free at high shutter speeds.

- Two side-angle lights (~45°) and one fill light (optional).

- Enables short exposures (1/500–1/1000 s) at 60 fps for motion-blur-free imaging.

8. Expected Performance

- 60 fps recording per camera at 1456 × 1088 px.

- Synchronization precision: within microseconds.

- Global shutter eliminates rolling distortion.

- High contrast, sharp, low-blur video suitable for clinical kinematic analysis.

9. Next Development Steps

- Assemble the three Pi + camera units with lenses, cooling, and tripods.

- Install Raspberry Pi OS and libcamera; confirm stable 60 fps capture.

- Wire the external trigger between Pis (one master, two slaves).

- Calibrate lighting and verify correct exposure (~1/750 s).

- Validate synchronization via a shared LED flash visible to all cameras.

- Record trial gait videos; review in Kinovea for alignment and clarity.