Alyn Hospital

Gait Visualization System

User manual

Introduction

There is a need for developing a video system which will help mechanotherapists to analyze the gait of young patients in Alyn Children Rehabilitation hospital.

The Gait Visualization System is a video system based on multiple cameras, to be used in a dedicated room. The patient is instructed to walk a few meters in the room, while a set of synchronized cameras records his motion through the room. The resulting video clips are then simultaneously displayed on a common computer screen, so the mechanotherapists can review and diagnose the details of the patient's body movements.

General Architecture

The proposed system consists of at least three video cameras (shutter visual sensors). Two of the cameras are lateral, one is frontal. Each camera is controlled by a single-board computer (SBC) RaspberryPi 4. Each camera records a single video stream and stores it in local storage of the corresponding SBC. The cameras are interconnected through a standard Ethernet network. A dedicated PC reads the video streams and displays them contiguously on screen.

See the diagram below:

Functionality

There are 4 functionality modes.

- Setup Mode

In this mode, the Master Pc is used to develop and upload software into the RaspberryPi SBCs. The communication is through standard tools such as RealVNC for remote control and FileZilla for file upload/download.

- Initialization Mode

In Initialization mode, the cameras are activated and the resulting video streams are forwarded in real time to the Master PC and displayed on the screen. This allows physical alignment of the cameras to the correct position and tilt, pan.

- Recording Mode

In Recording mode, the Master PC sends a command to the SBCs, consisting of start time and duration of the recording (assuming SBCs and Master PC are synchronized using RTP). Following the command, the cameras are simultaneously activated, and the resulting video streams are stored into local memory (SD card, or SSD). Another command can be used to stop recording.

- Analyze mode

In Analyze mode, the recorded streams are uploaded to the Master PC for analysis. This can be done through storage sharing between the SBCs and Master Pc.

- Setup

5.1	OS

Install OS on SD card to obtain:

PRETTY_NAME="Debian GNU/Linux 13 (trixie)"

NAME="Debian GNU/Linux"

VERSION_ID="13"

VERSION="13 (trixie)"

Then upgrade all:

sudo apt upgrade

Have one user:

Username: alyn

Password: alyn

  - SMB

Note: samba is not reliable. We use ssh instead.

Install Samba:

sudo apt install samba samba-common-bin -y

make a shared directory:

mkdir videos

modify /etc/samba/smb.conf, by adding at the end:

[videos]

comment = Raspberry Pi Shared Folder

path = /home/alyn/videos

browseable = yes

writeable = yes

only guest = no

create mask = 0777

directory mask = 0777

public = no

valid users = alyn

force user = alyn

then

sudo smbpasswd -a alyn

sudo systemctl restart smbd

sudo systemctl status smbd

Note:

On PC, login to each rpi in Windows browser, by login to:

\\192.168.50.11(/12/13/14)\videos: user – alyn, password: alyn, and keep login got later use.

  - LAN setting

Use Tftpd64 tool to set DHCP as follows (FIXME)

  - RPI configuration

Disable WAN and enable ethernet as follows:

Assign fixed and known IP to each Raspberry Pi ethernet:

sudo nmtui

Edit a connection
→ Wired connection 1

IPv4 CONFIGURATION
→ Manual

| Field | Value |
| --- | --- |
| Address | 192.168.50.11(or 12 or 13 or 14 respectively)/24 |
| Gateway | empty |

Optional DNS:

8.8.8.8

Then on PC:

Login to each rpi in Windows browser, by login to:

\\192.168.50.11(/12/13/14)\videos: user – alyn, password: alyn, and keep login got later use.

5.5 Raspberry Pi Synchronization

Install chrony on rpi:

sudo apt install chrony -y

sudo nano /etc/chrony/chrony.conf

Add near top:

server 192.168.50.1 iburst

Enable NTP server on PC:

On PowerShell:

Set-Service w32time -StartupType Automatic

Start-Service w32time

On Admin CMD:

reg add HKLM\SYSTEM\CurrentControlSet\Services\W32Time\TimeProviders\NtpServer /v Enabled /t REG_DWORD /d 1 /f

reg add HKLM\SYSTEM\CurrentControlSet\Services\W32Time\Config /v AnnounceFlags /t REG_DWORD /d 5 /f

on PowerShell:

New-NetFirewallRule -DisplayName "NTP Server UDP 123" -Direction Inbound -Protocol UDP -LocalPort 123 -Action Allow -Profile Private

On admin CMD:

Restart-Service w32time

w32tm /resync

Install synchronization service

On RPI:

Synchronization is based on sync_clock.py script.

Sync_clock.py is activated at boot by

/etc/systemd/system/syn_clock.service script with the following content:

[Unit]

Description=Synchronize Pi clock with Windows PC

Wants=network-online.target

After=network-online.target

[Service]

Type=simple

User=root

ExecStart=/usr/bin/python3 -u /home/alyn/gait/sync_clock.py

Restart=always

RestartSec=5

[Install]

WantedBy=multi-user.target

To run the service type the following:

sudo systemctl daemon-reload

sudo systemctl enable --now sync-clock.service

On PC

First go to gait directory

cd C:\Users\<user>\OneDrive\מסמכים\Alyn\gait analysis

Then run the sync service:

python3 time_server.py

Install ffmpeg in PC

Winget install ffmpeg

Automatic Rpi start

Create a service named gait-video:

sudo nano /etc/systemd/system/gait-video.service

The content of the file is:

[Unit]

Description=gait camera Service

After=network.target

[Service]

ExecStart=/usr/bin/python3 /home/alyn/gait/capture_StreamVideo.py

WorkingDirectory=/home/alyn/gait

StandardOutput=inherit

StandardError=inherit

Restart=always

User=alyn

[Install]

WantedBy=multi-user.target

Then run:

sudo systemctl daemon-reload

sudo systemctl enable gait-video.service

sudo systemctl start gait-video.service

SSH initialization

On Windows

Generate a key:

ssh-keygen

Press Enter for all prompts.

This creates:

C:\Users\<username>\.ssh\id_rsa
C:\Users\<username>\.ssh\id_rsa.pub

(or id_ed25519 on newer systems)

Copy the public key to the Pi

Run:

ssh-copy-id alyn@192.168.1.101

If ssh-copy-id isn't available on Windows:

type $env:USERPROFILE\.ssh\id_ed25519.pub

Copy the contents.

Then on the Pi:

mkdir -p ~/.ssh
nano ~/.ssh/authorized_keys

Paste the key and save.

Set permissions:

chmod 700 ~/.ssh
chmod 600 ~/.ssh/authorized_keys

Now test:

ssh alyn@192.168.1.101

You should log in without a password.

Then:

scp alyn@192.168.1.101:/home/alyn/videos/video.mp4 .

will also work without a password.

You can use the same key pair for all of your Raspberry Pis, and that's actually the most common approach.

Recommended setup

On your Windows PC:

ssh-keygen -t ed25519

This creates:

C:\Users\<your_user>\.ssh\id_ed25519
C:\Users\<your_user>\.ssh\id_ed25519.pub

Then copy the same public key (id_ed25519.pub) into:

/home/alyn/.ssh/authorized_keys

on each Raspberry Pi.

Result:

Windows PC
    │
    ├── SSH key
    │
    ├── RPi #1
    ├── RPi #2
    ├── RPi #3
    └── RPi #4

One key can authenticate to all of them.

4. Check SSH server config on Pi

sudo grep -E "PubkeyAuthentication|AuthorizedKeysFile|PasswordAuthentication" /etc/ssh/sshd_config

You want at least:

PubkeyAuthentication yes
AuthorizedKeysFile .ssh/authorized_keys

Then restart SSH:

sudo systemctl restart ssh

Most common fix:

chmod 700 /home/alyn/.ssh
chmod 600 /home/alyn/.ssh/authorized_keys
chown -R alyn:alyn /home/alyn/.ssh

Improvements

- Guidance for users: כיול, גובה המצלמות, זווית, פוקוס, זום

- סווג של סרטונים עם מס' תיק, תאריך, מצב מתועד

- Start/stop for capture

- Automatic display

- Serial number per file

- Folder for date, subfolder for file number

- FPS = 50

Notes

- Rpi1 is replaced by rpi5

- NTP source of rpi5 is rpi4, not PC.

Improvements

- Frame rate: is it correct?

- GUI